STATUS:  Passed  Service Required  Preventable problem  Adverse event likely

Defibrillators and Noninvasive Pacemakers

Procedure No. 408-20081015-01 ( Major )

| Facility: | Dept: |  |
| --- | --- | --- |
| WO No.: | Date: | Inspector: |
| Mfr: | Model: | SN: |
| CTRL No: | Loc: | IPM Time: |

SPECIAL PRECAUTIONS

WARNING:During defibrillator discharge or pacing, the high voltage present at the paddles or electrodes is extremely dangerous and possibly lethal. Never perform tasks alone. A second person must be present to summon help and/or apply CPR in the event of an emergency. Never apply the defibrillation or pacing paddles or electrodes to anyone, and never hold or contact the conductive electrode portion of the defibrillation or pacing electrodes or paddles. Testing input isolation requires the use of a line voltage source. Although this source should include a current-limiting resistor, use caution to avoid contact with any portion of the circuit while it is energized. Defibrillators and pacemakers must always be available to clinical staff in the event of an emergency and inspection testing may deplete the battery. Either the inspection should be performed in the vicinity of the defibrillator's usual storage location with a fully charged battery available before beginning testing or a unit with which the clinical staff is familiar should be made available as a substitute during the inspection. Do not test all the units in an area at one time, as this will leave the staff inadequately equipped to handle emergencies.

| TEST APPARATUS |
| --- |
|  | Control/Serial No. |
| Electrical safety analyzer (may be included in defibrillator analyzer) |  |
| Defibrillator analyzer with ECG simulator (VF and normal sinus rhythm waveforms) |  |
| Stopwatch or watch with a second hand |  |
| Noninvasive pacemaker analyzer (for testing units with pacing capability and if not included in defibrillator analyzer) |  |
| Adapters for connecting pacing cable leads to test equipment (for units with pacing capability) |  |

| QUALITATIVE TASKS |
| --- |
| Pass | Fail |  |
|  |  | Chassis/Housing |
|  |  | Mounts/Fasteners |
|  |  | Casters/Brakes |
|  |  | AC Plug |
|  |  | Receptacles |
|  |  | Line Cord |
|  |  | Strain Reliefs |
|  |  | Cables |
|  |  | Fittings/Connectors |
|  |  | Paddles/Electrodes |
|  |  | Controls/Switches |
|  |  | Recorder (ECG) |
|  |  | Battery/Charger |
|  |  | Indicators/Displays |
|  |  | Self-Test |
|  |  | Time/Date Settings |
|  |  | Network/Wireless Interfaces |
|  |  | Synchronizer |
|  |  | Integral Output Tester |
|  |  | Alarms |
|  |  | Pacemaker Demand-Mode Activation/Inhibition |
|  |  | Audible Signals |
|  |  | Labeling |
|  |  | Accessories |

| QUANTITATIVE TASKS |
| --- |
|  | Criteria | Set/Indicated | Measured | Pass | Fail |
| Grounding Resistance | ≤0.5 Ω |  |  |  |  |
| Manual Mode Defibrillator Output | ±4 J low; ±15% high |  |  |  |  |
| Energy after 60 Sec (Manual Mode) | ≥85% |  |  |  |  |
| Heart Rate | ±5% |  |  |  |  |
| Heart Rate Alarm |  |  |  |  |  |
| Pacing Rate | ±5% |  |  |  |  |
| Pacing Amplitude | ±10% |  |  |  |  |
| Automated Mode Analysis and Defibrillator Output |  |  |  |  |  |
| Multiple Discharge Output Energy and Charge Time | ±4 J low; ±15% high |  |  |  |  |
| Pediatric Mode Output Energy | ≤50 J |  |  |  |  |
| Internal Paddle Energy Limit | ≤50 J |  |  |  |  |

| PREVENTIVE MAINTENANCE |
| --- |
| Done |  |
|  | Replace |

NOTES