General-Purpose Infusion Pumps

Procedure No. 416-20081015-01 ( Major )

Used For:

Infusion Pumps, General-Purpose  [13-215]

Commonly Used In:

All patient care areas, homes

Scope:

Applies to general-purpose infusion pumps

Risk Level: High

| Type | Interval | Time Required |
| --- | --- | --- |
| Major | 12 Months | 0 hours |
| Minor | 0 NA | 0 hours |

Overview:

Infusion pumps are often used when accurate delivery rates are required over long periods of time. Most general-purpose infusion pumps have a flow range of 0.1 to 999 mL/hr.  Some general-purpose infusion pumps are multichannel models, offering two or more general-purpose pumps within one chassis. Because general-purpose pumps now offer flow resolution and flow settings down to 0.1 ml/hr, they are now commonly used in neonatal critical care areas.

Test Apparatus, Supplies, Parts:

- Electrical Safety Analyzer

- Pressure meter (≥50 psi)

- Infusion pump tester - optional

- 50 mL graduated cylinder (1 mL graduations)

- Electronic balance (200 g range, 0.1 mg resolution)

- 5 mL pipette (0.1 mL graduations)

- Stopwatch or watch that displays seconds

Special Precautions:

Procedure:

Be sure that you understand how to operate the equipment, the significance of each control and indicator, and the alarm capabilities. Before beginning an inspection, carefully read this procedure, the operators manual, and the inspection and preventive maintenance procedures recommended by the manufacturer (typically included in the service manual). Use the BiomedicalBenchmark Support Assessment Form to document a maintenance decision that reflects past experience with this type of equipment and the environment where it is used. Then use the IPM Procedure Customization Tool to modify this procedure as needed; the program will generate a documentation form with the corresponding changes. Also see: Inspecting Infusion Pumps: How Often is Enough? (in IPM Guidance article Optimizing the IPM Program).

Use the disposable accessory (e.g., IV tubing set) specified for the infusion pump to be inspected. Disposable accessories should be replaced at least weekly and more frequently with heavy use. Use degassed water or, when available, bags of outdated (i.e., clinically unusable) IV solutions.

For settings ≥1 mL/hr, flow accuracy is most conveniently determined with an infusion pump tester. If a tester is not available, calculate flow rate by dividing the volume collected in a 50 mL graduated cylinder (1 mL graduations) by the delivery time measured with a stopwatch. For settings less than 1 mL/hr, calculate flow rate by dividing the mass obtained in a small beaker on an electronic balance (200 mg range and resolution to 0.1 mg) by the measurement time. If a balance is not available, use a 5 mL pipette (0.1 mL graduations) attached to a vertical mounting stand and divide the volume collected by the delivery time.

Operate the pump on battery power during during a flow accuracy measurement if not during the entire inspection to check that the battery has been charged and can hold a charge. If a low-battery alarm occurs, verify that the alarm is properly displayed and then continue the inspection on line power. Note how long the pump was operating on battery power and the conditions under which the low-battery alarm occurred. Fully charge the battery before returning the pump to use.

Qualitative tasks: Chassis/Housing.

Examine the infusion pump for overall condition. The chassis should be clean and free from IV or enteral solution residue, especially near moving parts (e.g., thumbwheel switches, pump or controller mechanisms). Also check for dried solution deposits on accessible air-in-line sensors, pressure sensing mechanisms, and infusion set/cassette locking mechanisms. Check that labels and markings are legible.

Mount.

Screws and brackets that attach the infusion pump to an IV pole should be secure and functioning. If the pump is mounted on a stand or cart, examine the condition of the mount. Also examine the pole, stand, or cart.

Casters/Brakes.

If the infusion pump is mounted on a dedicated IV pole, stand, or cart that moves on casters, check their condition. Look for accumulations of lint and thread around the casters and be sure that they turn and swivel, as appropriate. Check the operation of brakes and swivel locks, if so equipped.

AC Plug.

Examine the AC power plug for damage. Attempt to wiggle the blades to determine that they are secure. Shake the plug and listen for rattles that could indicate loose screws. If any damage is suspected, open the plug and inspect it.

If the pump or its IV pole has electrical receptacles for accessories, inspect them by inserting an AC plug into each and checking that it is held firmly. If accessories are plugged and unplugged often, consider a full inspection of the receptacle.

Line Cord.

Inspect the cord for signs of damage. If damaged, either replace the entire cord or, if the damage is near one end, cut out the defective portion.

Strain Reliefs.

Examine the strain reliefs at both ends of the line cord. Be sure that they hold the cord securely.

Fittings/Connectors.

Examine any electrical cable connectors (e.g., data transfer, nurse call) for general condition. Electrical contact pins or surfaces should be straight and clean. Check any spill-protection connector caps for signs of damage.

Controls/Switches .

Examine all controls and switches for physical condition, secure mounting, and correct motion. Where a control should operate against fixed-limit stops, check for proper alignment, as well as positive stopping. Check membrane switches for membrane damage (e.g., from fingernails, pens). During the course of the inspection, be sure to check that each control and switch performs its proper function. Verify operation of the front keypad lockout switch, located on the back of most models.

Battery.

If so equipped, check the pump's battery log for battery status. Operate the pump on battery power during during a flow accuracy measurement if not during the entire inspection to check that the battery has been charged and can hold a charge. If a low-battery alarm occurs, check to ensure that it is properly displayed and then continue the inspection using line power. Note how long the infusion pump has been operating and the conditions under which the low-battery alarm occurred. Fully charge the battery before returning the pump to use. When it is necessary to replace a battery, label it with the date.

Indicators/Displays.

During the course of the inspection, confirm the operation of all lights, indicators, meters, gauges, visual displays, and display backlighting, if so equipped. Be sure that all segments of a digital function. (Many infusion pumps automatically check indicator and display function when turned on or during a manually activated self-test.)

Self-Test.

For units with a self-test mode, activate it and determine if the expected response is produced.

Time/Date Settings.

Verify that the time and date settings on the unit are correct.

Network/Wireless Interfaces.

Review measures taken to ensure protection against the potential risks and vulnerabilities to the confidentiality, integrity, and availability of electronic information stored or transmitted by the device or system and verify that preventive measures are still active. For example, are passwords being applied correctly, have operating system and virus protection patches and upgrades been installed, is the device still operating on a VPN (virtual private network) and are wireless security measures still in place. Verify that data backup processes are activated and that data can be retrieved from backups.

Alarms.

Many infusion pump alarm capabilities can be checked qualitatively. The following tasks include tests for the most common alarm conditions. Check the instruction manual to see how the alarm should work. When an alarm occurs, check to see that both audible and visual alarms are activated and that flow stops or is reduced to a keep-vein-open rate (e.g., <5 mL/hr). Confirm appropriate alarm volume, as well as the operation of any volume control.

Audible Signals.

Operate the pump (e.g., press rate switches) to activate any audible signals. Confirm appropriate volume, as well as the operation of the volume control, if so equipped.

Open Door/Misloaded Infusion Set.

Check this alarm during setup and operation.

Flow-Stop Mechanism(s).

Turn the power off with the infusion set primed and loaded in the device. With all tubing clamps open, the fluid container two feet or more above the device, and the tubing hanging to the floor, verify that no fluid flows out of the distal connector.

If the device incorporates a mechanism that automatically closes the set or requires the set to be manually closed before it is removed from the device, verify the operation of this mechanism.

Occlusion.

Verify upstream occlusion detection by activating infusion with the tubing clamped just below the fluid container. Verify operation of downstream occlusion detection by activating infusion with the infusion set's distal connector capped.

Nurse Call (verify only if this function is used).

Some pumps have a relay contact closure that activates a nurse-call system when an alarm condition occurs. This requires a special cable that connects the pump to the nurse-call system. If the infusion pump has this capability and it is used in any clinical location, connect the cable, and simulate one or more of the above alarm conditions to determine whether they activate the nurse call. Alternatively, use an ohmmeter to check that a change in resistance (either low to high or high to low) occurs between the two conductors of the cable when an alarm condition is created.

Labeling.

Check that all necessary placards, labels, and instruction cards are present and legible.

Quantitative tasks: Grounding Resistance.

Measure the resistance between the grounding pin of the power cord (if so equipped) and exposed (unpainted and not anodized) metal on the chassis. Grounding resistance should not exceed 0.5 Ω. If the unit is double insulated, grounding resistance need not be measured.

Flow Accuracy.

Determine the flow accuracy at two typical flow settings ≥1 mL/hr (e.g., 10 and 100 mL/hr).

Use an infusion pump tester or collect the output in a graduated cylinder. Use a stopwatch or a watch that displays seconds to time the delivery into the graduated cylinder until at least 10 mL is collected. Record the time interval and volume collected, and calculate the delivery rate in mL/hr.

To calculate flow error, use the following formula:

Most infusion pumps are specified to deliver within 5% of the flow setting.

Note: Negative and positive flow error represents underdelivery and overdelivery, respectively.

Maximum Pressure.

Reference the pump's maximum pressure determined during acceptance testing. Connect the distal end of the primed administration set to a pressure meter or use an infusion pump analyzer (if equipped to perform this task). Start infusion at a commonly used flow setting (e.g., 10 or 100 mL/hr) and record the maximum pressure. Restart the pump after an occlusion alarm to ensure that maximum infusion pressure has been attained. The maximum pressure of most pumps is less than 20 psi.

Air-in-Line Detection.

Inject 100 μL air into an injection port of the IV tubing with a U-100 insulin syringe between the fluid container and the air-in-line detector. 100 μL can be approximated by 10 units from a U-100 insulin syringe. Sensitivity to air volumes of less than 50 μL is likely to result in nuisance alarms. Most pumps will trigger an alarm for 50 to 100 μL air.

Preventive Maintenance: Replace.

Replace the primary battery and/or the clock/memory battery/cell, if necessary, and label it with the date.