Nd-YAG Surgical Lasers

Procedure No. 447-20081015-01 ( Major )

Used For:

Lasers, Nd:YAG, Surgical [16-943]

Commonly Used In:

Operating rooms, short procedure areas, cystology rooms, catheterization laboratories, endoscopy laboratories, radiology areas

Scope:

Applies to general-purpose Nd:YAG surgical lasers that include contact (e.g., sapphire tip) and/or noncontact flexible fiberoptic delivery systems (either reusable or disposable), emit near-infrared energy at 1,064 nm, and can provide sufficient power output to coagulate and vaporize tissue; applies to low- and high-power Nd:YAG surgical lasers that are typically used for general surgery, urology, cardiovascular surgery, gastroenterology, bronchopulmonary, neurosurgery, gynecology, ENT, and plastic surgery procedures; applies to the Nd:YAG portion of units that combine the Nd:YAG wavelength with other wavelengths (e.g., KTP or CO2); does not apply to ophthalmic Nd:YAG lasers or to Nd:YAG lasers that are, for example, frequency-doubled and do not emit energy for delivery to the patient at 1,064 nm (frequency-doubled lasers are covered in Procedure 464); also does not apply to CO2 lasers, argon lasers, or other ophthalmic lasers; however, many of the tests listed herein can be used or modified for these other lasers

Risk Level: High

| Type | Interval | Time Required |
| --- | --- | --- |
| Major | 12 Months | 0 hours |
| Minor | 6 Months | 0 hours |

Overview:

Nd:YAG lasers are normally checked before each use by the laser's power-on self-test and by user examination of the aiming beam and calibration of the system with the delivery system to be used. This minimizes the need for frequent additional periodic testing.

Manufacturers or outside service vendors often maintain lasers for hospitals. The extent and frequency of inspection by hospital personnel should be coordinated with these outside services.

Failure of an Nd:YAG surgical laser can cause patient or staff injury, abrupt interruption of a surgical procedure, or damage to the laser system. These lasers must be meticulously maintained to ensure proper and safe operation.

Nd:YAG surgical lasers affect tissue by delivering invisible near-infrared energy at a sufficient power density to cause photocoagulation, thermal denaturation, and/or vaporization of tissue. The 1,064 nm Nd:YAG energy is not well absorbed by any tissue and is typically scattered over a 5 mm depth within tissue. Nd:YAG surgical lasers are frequently used to cause photocoagulation or thermal denaturation of tissue in the noncontact mode or with quartz or sapphire contact tips.

General-purpose Nd:YAG surgical lasers have a laser cavity that houses an yttrium-aluminum-garnet (YAG) crystalline rod that is doped with neodymium (Nd). Energy leaving the laser tube through a partially reflecting mirror is typically directed into a flexible optical fiber that transmits the laser energy to the tissue. The fiber may be used with additional connectors (e.g., through an endoscope), with contact tips or contact-tip fibers connected, and/or with a laser handpiece or a laser micromanipulator (used to interface the laser with the surgical microscope) connected. These attachments can focus the energy into a small spot size at a known working distance or a specific beam pattern to accomplish a specific task. In a few cases (e.g., the Trimedyne Laserprobes SLT contact tips), the laser energy is transformed into thermal energy to heat a catheter tip, which then causes the clinical effect; in this case, no or very little direct laser irradiation of tissue occurs.

Because the near-infrared energy emitted by the Nd:YAG laser is invisible, a second, nontherapeutic aiming helium-neon (He-Ne) laser—which emits visible red light—a red diode laser, or a xenon lamp with filters to emit white or blue light simultaneously traverses the fiber and is coincident (i.e., travels the same path) with the Nd:YAG laser beam.

Like most lasers, Nd:YAG lasers are inefficient in converting electrical energy into laser energy of 0 to 140 W. As a result, excess heat is generated in the laser tube, requiring a cooling system. Some Nd:YAG lasers use internal water/air cooling systems, while others require external connection to a water source and drain or to a freestanding cooling system. Most Nd:YAG laser fibers require gas or liquid cooling as well for certain applications. As a result, gas compressors/regulators and/or fluid pumps are typically integral or attached to these lasers. Because Nd:YAG laser fibers may be used in body cavities (e.g., during gastroscopy), some lasers may include a gas recirculation system that inserts gas to cool the fiber and/or insufflate the cavity and withdraw gas to limit pressure and avoid distension.

With Nd:YAG lasers, unlike those lasers that use mirror delivery systems (e.g., articulating arms on CO2 lasers), it is not necessary to periodically verify coincidence of the aiming and therapeutic beam or to assess the therapeutic beam pattern within the beam or spot. Since the therapeutic and aiming laser beams are transmitted through a single optical fiber, these two beams are coincident as they exit the fiber. Any beam pattern distortion at the fiber entrance would be eliminated as the laser beams travel through the fiber because of internal reflections within the fiber. Misalignment of the beam at the fiber entrance would result in decreased power output or loss or distortion of the aiming beam. In a well-aligned system, any significant problem with the therapeutic beam pattern introduced by an accessory would be apparent by examining the visible aiming beam.

Test Apparatus, Supplies, Parts:

- Electrical safety analyzer

- Laser radiometer (0.01 to 20 W, 0.1 to 100 W, 0.01 to 10 sec)

- Calibrated flowmeter (optional)

- Outlet test fixture (optional)

Special Precautions:

Inspecting and maintaining lasers is a dangerous but necessary process, and far greater care is required than with most devices. Personnel who inspect or service lasers should receive special training from the manufacturer or from a qualified alternative training source.

Laser energy can cause serious injury, particularly when the internal interlock is overridden or in any other situation in which the energy does not diverge significantly over long distances. Under some circumstances, the beam may not diverge significantly, even a full room length or more away from the laser (and can harm tissue or burn material even at this distance). Therefore, exercise great care whenever a laser beam is accessible. Area security and use of personnel protective devices and practices should be consistent with hospitalwide laser safety procedures and/or be approved by the laser safety committee.

In addition, windows should be covered with laser-opaque, nonreflective material to prevent transmission of laser energy to other areas.

Wear appropriate laser safety eyewear at all times whenever the laser is in the operating mode. WARNING: Laser safety eyewear may not protect the wearer from the aiming system light. Do not stare directly into the aiming system beam or the therapeutic laser beam, even when wearing laser safety eyewear. Avoid placing the laser beam path at eye level (i.e., when kneeling, sitting, or standing).

Do not perform these procedures when a patient is present or clinical staff are working, and do not aim the laser across a path that a person might normally use as a thoroughfare. Furthermore, at minimum, post doors to the room with appropriate laser safety signs stating that the laser is in use and that it is unsafe to enter the room without authorization by the service person performing the procedure. A second person should be present, especially during procedures of recognized risk, to summon help in case of an accident.

The laser should remain in the off position when not in use. When in use, the laser should be in the standby/disabled mode. Do not switch it to the operating mode until the procedure is about to begin and the laser and its delivery system are properly positioned. If the procedure must be interrupted, disconnect the laser from line voltage, and remove the laser operation key and store it in a controlled location.

Do not use the laser in the presence of flammable anesthetics or other volatile substances or materials (e.g., alcohol), or in an oxygen-enriched atmosphere, because of the serious risk of explosion and fire. Remove from the working area or cover with flame-resistant opaque material all reflective surfaces likely to be contacted by the laser beam. Whenever possible, use a firebrick or other nonflammable material behind the target material (e.g., black Delrin) when the laser is to be activated. A CO2 fire extinguisher should be readily available.

Some surgical lasers use high voltages (e.g., 20 kV), which can be lethal. Capacitors may store charges long after the device has been disconnected from line voltage. Consult the manufacturer's recommended procedures for servicing high-voltage laser circuits, and avoid contact with any portion of the high-voltage circuit until you are certain that the charge has been drained. In such instances, a good ground must be present; preferably, use a redundant ground strap if you must enter the laser cabinet. When possible, disconnect the laser from line voltage before entering the laser cabinet, and use insulated gloves for those procedures in which contact with a high-voltage source is possible (and the gloves are not otherwise contraindicated). Ensure that equipment intended to be used to measure, drain, or insulate high voltages carries the appropriate insulation rating (e.g., above 20 kV).

Where possible, perform tests with the laser turned off. Because of the presence of high voltage, perform the Grounding Resistance test before any other task that requires operation of the laser.

WARNING: Lasers may be damaged by switching between normal and reverse polarity while the device is on. If reverse-polarity leakage current measurements are made, turn off the laser being tested before switching polarity. Also, lasers powered by three-phase electrical systems may be damaged if proper electrical phase connections are not made initially and maintained thereafter. Thus, do not switch conductor connections or wiring configurations for any tests, including leakage current measurement. Do not conduct electrical leakage current tests with reversed-polarity wiring.

WARNING: Do not use an anesthesia or other similar bag that may have a mold-release agent (e.g., starch, talc) on its inside surface because the agent could contaminate the gas recirculation system of the laser and ultimately contaminate a patient wound during a subsequent procedure.

WARNING: In general, do not use liquid fiber cooling for tests unless specifically described in the task. Use of this kind of cooling rather than gas fiber cooling may damage test equipment or cause erroneous test results.

Report any laser accident immediately to the laser safety officer or equivalent, as well as to the hospital risk manager.

Procedure:

Be sure that you understand how to operate the equipment, the significance of each control and indicator, and the alarm capabilities. Before beginning an inspection, carefully read this procedure, the operators manual, and the inspection and preventive maintenance procedures recommended by the manufacturer (typically included in the service manual). Use the BiomedicalBenchmark Support Assessment Form to document a maintenance decision that reflects past experience with this type of equipment and the environment where it is used. Then use the IPM Procedure Customization Tool to modify this procedure as needed; the program will generate a documentation form with the corresponding changes.

Qualitative tasks: Chassis/Housing.

Verify that the key has not been left in the laser. (Remove it if it has, and inform users of the importance of storing the key in a controlled location.) Examine any external gas tanks that may be in use with the laser, and ensure that they have been turned off after the last use. Examine the exterior of the laser for cleanliness and general physical condition. Be sure that all housings are intact and properly aligned, that assembly hardware is present and tight, that any retractable parts slide easily and lock in place if so constructed, that there are no signs of spilled liquids or other evidence of abuse, and that there are no obvious signs of water or oil leakage.

Shutters.

If manual shutters for the aiming system or therapeutic laser are accessible, ensure that they operate smoothly and correctly. Be sure to leave the shutter in the proper position for normal operation.

Mounts/Holders.

Ensure that the assembly is stable and that the laser will not tip over when pushed or when a caster is jammed on an obstacle (e.g., a line cord, threshold), as may occur during transport. If the laser is mounted on a stand or a cart, examine the condition of the mount. Verify that the mounting apparatus is secure and that all hardware is firmly in place.

AC Plug.

Examine the AC power plug for damage. Wiggle the blades to determine whether they are secure. Shake the plug, and listen for rattles that could indicate loose screws. If you suspect damage, open the plug and inspect it.

Line Cords.

Inspect line cords for signs of damage. If a cord is damaged, replace the entire cord, or, if the damage is very near one end, cut out the defective portion.

Strain Reliefs.

Examine the strain reliefs at both ends of the line cord. Be sure that they grip the cord securely.

Circuit Breakers/Fuses.

If the laser has a switch-type circuit breaker, check that it moves freely. If the laser is protected by an external fuse(s), check its value and type against what is marked on the chassis or noted in the instruction or service manual. Ensure that a spare is provided or readily available.

Tubes/Hoses.

Check the condition of all cooling-system hoses and any other hoses or tubing the laser may have (e.g., drain, gas). Check that they are of the correct type; that they have not become cracked and do not show other signs of significant abuse; that they are connected correctly and positioned so that they will not leak, kink, trail on the floor, or be caught in moving parts; and that they are secured adequately to any connectors.

Cables.

Inspect all cables and their channels or strain reliefs for general physical condition. Examine cables carefully to detect breaks in insulation and to ensure that they are gripped securely in the connectors at each end to prevent strain on the cable.

Fittings/Connectors.

Examine all optical (e.g., fiber), gas, liquid, and electrical fittings and connectors for general physical condition. Gas and liquid fittings should be tight and not leak. Electrical contacts should be straight, clean, and bright.

If external gas tanks or wall-supply outlets can be used, gas-specific connectors should be present. Be sure that no pins are missing from yokes and that the keying and indexing of connectors for each gas to be used is correct. A laser that connects to a central piped medical gas system or to a freestanding medical gas system should have the matching DISS or quick-connect fitting for the gas that it is to be used with. Verify that suitable connectors are supplied so that adapters are not required.

Filters.

Check the condition of all liquid and air filters. Some Nd:YAG surgical lasers require deionized water, and most require special filtration. Measuring the pressure drop across a liquid filter can be helpful in determining whether the filter should be replaced. Clean or replace filters according to the manufacturer's recommendations (e.g., replace if the pressure drop is >5 psi), and indicate this in the preventive maintenance section of the inspection form. Clean or replace air filters and radiators that are obviously dirty.

Controls/Switches.

Before moving any controls, check and record their positions. If any position appears unusual, consider the possibility of inappropriate use or of incipient device failure. Examine all controls and switches for physical condition, secure mounting, and correct motion. If a control has fixed-limit stops, check for proper alignment, as well as positive stopping. Check membrane switches for tape residue and for membrane damage (e.g., from fingernails, pens, surgical instruments). If you find such evidence, notify users to avoid using tape and sharp instruments. During the inspection, be sure that each control and switch works properly.

Remote.

Examine the exterior of the control for cleanliness and general physical condition. Be sure that housings are intact, that assembly hardware is present and tight, and that there are no signs of spilled liquids or other serious abuse. If the remote control is attached by cable to the laser, ensure that the cable and any connectors are in good condition. Examine all controls and switches for general physical condition, secure mounting, correct motion, and intended range of settings. Where a control should operate against fixed-limit stops, check for proper alignment, as well as positive stopping. During the inspection, be sure to check that each control and switch performs properly.

Footswitch.

Examine the footswitch for general physical condition, including evidence of spilled liquids. Footswitches for lasers include an internal switch that activates according to the depth of pedal depression. It is usually possible to feel the vibration caused by closure of the switch, even through a shoe. Check that the internal switch is operating and that the footswitch does not stick in the on position. Some footswitches include two internal switches; in this case, verify the operation of both. Some footswitches also include a switch to operate the liquid- or gas-cooling system. Check to be sure that this switch operates reliably.

During the procedure, check to be sure that the laser activates and deactivates consistently when the footswitch is depressed and that the fiber-coolant system operates properly when the fiber-coolant switch is activated and deactivated. Flex the cable at the entry to the switch, and, using an ohmmeter, check for internal wire breaks that cause intermittent operation. Confirm that strain reliefs are secure.

Examine the male and female connectors for attaching the footswitch to the laser cabinet to be sure that no pins are bent and that no other damage is present. Ensure that the connector secures acceptably to the laser cabinet.

Pumps/Fans/Compressors.

Check the physical condition and proper operation of these components, if present. If lubrication is required, note this in the preventive maintenance section of the inspection form. Clean any obvious dust from these components.

Fluid Levels.

Check all fluid (e.g., coolant) levels. Refill or change the fluid according to the manufacturer's recommendations, and note this in the preventive maintenance section of the inspection form. If an external water supply is in use, ensure that the water pressure is properly regulated and at the appropriate pressure and that the supply and drain system is properly configured (e.g., filters are oriented for proper flow, drain hoses are positioned in a sink or drain).

Battery.

If a remote control or display is battery powered, check or replace the battery (periodic prophylactic battery replacement is often preferred to risking battery failure during use). Other batteries, such as those used to provide additional power for 110 VAC Nd:YAG lasers, should be inspected according to the manufacturer's recommendations. When it is necessary to replace a battery, label it with the date.

Indicators/Displays.

During the inspection, verify proper operation of all lights, indicators, meters, gauges, and visual displays on the laser and remote control. Ensure that all segments of a digital display function. Note any error messages displayed during the power-on self-test.

If primary and remote-control indicators and displays can be used at the same time or if control can be switched from one to the other during a procedure, operate the laser in a way that will verify that the same information (e.g., settings, displays) is indicated on both control panels.

If display screens or digital displays are provided for user prompts or for viewing accumulated information (e.g., pulse or accumulated energy counter), ensure that each display provides the information expected. Ensure that user prompts occur in the proper sequence. Store some sample information, and verify that it is correct. If a feature to manually reset this information is available, ensure that it works.

Time/Date Settings.

Verify that the time and date settings on the unit are correct.

Alarms/Interlocks.

Operate the laser in a manner that will activate the self-check feature, if present, and verify that all visual and audible alarms activate according to the manufacturer's documentation. If no self-check feature is present, operate the laser in a manner that will activate each audible and visual alarm; be sure to test only those alarms that will not cause damage to the laser or present an unnecessary risk of laser beam exposure.

If a door or window interlock is used, ensure that it properly deactivates the laser. (Do not disassemble major parts of the laser to test internal interlocks.) After deactivating the laser and reclosing the door or window, check to be sure that the laser will restart. Be sure to check the interlocks in all locations where the laser is used. (For some lasers, the function of the interlocks can be checked using an ohmmeter.)

If the laser is equipped with an emergency "kill" switch, test this feature to be sure that it deactivates the laser and that the laser will subsequently restart.

Audible Signals.

Operate the laser to activate any audible signals (e.g., laser emission, setting change). Check for proper operation, and verify that the signal can be heard in the environment in which the laser will be used.

Labeling.

Check that all placards, labels, and instruction cards noted during acceptance testing are present and legible. Check to see that the instruction manual is kept with the laser or is readily available.

Accessories.

Verify that all necessary accessories are available and in good physical condition. Set up reusable accessories with the laser to ensure compatibility and proper functioning.

Checking all fibers or accessories during a single inspection and preventive maintenance procedure is unnecessary as long as accessories are routinely checked by the person(s) responsible for laser setup and operation. In addition, many of the accessories are sterile and would require resterilization before use, making the laser potentially unavailable. Be sure to check with the person responsible for scheduling the use of the laser before beginning the procedure.

Fibers.

For the test fiber and before each use, examine the aperture connector, cable, and tip of each fiber to be used, as well as the fiber support, for cleanliness and general physical condition. Be sure that all hardware (e.g., laser gas tubing channels) is present, in good condition, and firmly attached. Ensure that the aperture connector properly seats into the laser aperture of the laser cabinet. Examine the distal end of fibers to ensure that any connecting mechanisms (e.g., threads) are in proper working order.

If a fiber appears to be dirty or damaged, remove it from service. If a fiber is reusable, notify the person(s) responsible for fiber repair. The fiber should be repaired and/or cleaned according to the manufacturer's recommendations. Verify fiber performance.

Contact Tips.

Examine each tip that may be used with the laser fibers for cleanliness and general physical condition. Be sure that the mechanism to connect a tip to a fiber is in proper working order and forms a secure connection. If a tip appears to be dirty or damaged, remove it from service and notify the person(s) responsible for tip repair or replacement. Some tips may look dirty after a single use, but remain acceptable for use; if you are unsure about the need to clean or repair a tip, consult with the person(s) responsible for tip repair or replacement and with the manufacturer, if necessary.

Handpieces.

Examine each handpiece component (e.g., body, tips, lenses) for cleanliness and general physical condition. Examine individually only those components that are intended for removal during normal use and storage. (Do not remove other parts that are press-fit or attached by screws, bolts, or snap-rings.) If lenses are detachable, be sure not to touch the lens surface; handle lenses by the edges only. Consult the manufacturer's recommendations for the procedures and cleaning agents to use to clean lenses.

Ensure that major subcomponents of the handpiece, when assembled, are secure. Ensure that the mechanisms used to connect the handpiece(s) to the fiber are in good working order and that they reliably secure each handpiece to the fiber.

Microscope Micromanipulator.

Examine the microscope micromanipulator for cleanliness and general physical condition. Be sure to handle it by the main body; do not hold it by the joystick, and do not touch the reflecting lenses in the body. Inspect micromanipulators provided by both the laser manufacturer and the laser accessory manufacturer.

Ensure that the reflecting surfaces and lenses are intact and clean. Consult the manufacturer's recommendations for the procedures and cleaning agents to use to clean reflecting surfaces and lenses.

Examine the joystick to ensure that it is firmly attached and that it freely moves the reflecting lens. If a finger rest is present, ensure that it is firmly attached and properly oriented.

If a zoom focus feature is present, be sure that it turns easily and does not slip. Examine each objective lens to ensure that it is intact and clean. Do not touch the lens surface. Consult the manufacturer's recommendations for the procedures and cleaning agents to use to clean the objective lenses. Carefully insert each lens into the micromanipulator, and ensure that it fits snugly.

There should be no visible dirt or residue in the optical path of the laser aperture. Ensure that any mechanism to close off the fiber laser aperture (fiber port) is clean, operates smoothly, and is in use.

Inspect the mechanism used to attach the micromanipulator to the microscope and the laser to ensure that all parts are present and that it is in good working order. Connect the micromanipulator to the microscope and the laser to check for a secure connection.

Safety Filters.

Verify operation of safety filters in microscope and endoscope delivery systems.

Aiming Beam.

Activate the aiming beam (without the therapeutic beam), and verify that it produces a round, uniformly bright spot, with no halo. For handpieces that provide adjustable spot sizes, verify that the spot size changes as expected and still remains uniform. Check that the intensity control, if present, does change the brightness of the aiming beam. Similarly, check pulsing controls to verify that the aiming beam can be pulsed. If several color choices are available for the aiming beam, verify that each color is present and working properly.

Laser Aperture.

WARNING: Make this inspection with the laser powered off.Remove and inspect the protective window (e.g., blast shield) behind the laser aperture. It should be clean and undamaged; clean or replace if needed. There should be no visible dirt or residue in the optical path of the laser aperture.

Gas Regulators.

Examine any gas regulators for cleanliness and general physical condition. Ensure that the gauges on the regulators are not broken. During the procedure, ensure that the regulator and the gauge operate as expected. Verify that the correct gas is attached to each regulator. Be sure that a key or wrench to facilitate changing the gas supply is with the laser or readily available.

If the laser includes a gas recirculation system, ensure proper operation by allowing it to control the gas supply into and out of a sealed plastic bag. If proper operation is questionable, consider using a calibrated flowmeter to measure actual gas flow.

Quantitative tasks: Grounding Resistance.

Measure the resistance between the grounding pin of the power cord (if so equipped) and exposed (unpainted and not anodized) metal on the chassis. Grounding resistance should not exceed 0.5 Ω. If the unit is double insulated, grounding resistance need not be measured.

Exposure Duration.

Some laser radiometers can measure pulse duration. If the power meter can react to pulse duration (this is the preferred circumstance), test the laser at each setting. However, if the radiometer does not measure pulse duration, use the following less preferable alternative.

Place and secure the laser fiber, handpiece, or micromanipulator with the aiming system focused on the target material (e.g., black Delrin or a tongue depressor). With the laser set to 10 W and the exposure set at minimum duration, activate the laser and create a burn. Carefully move the target material to expose a clean area, maintaining the same distance. Adjust the exposure setting in increments of 0.1 sec or the next longest duration, and activate the laser at each setting. Continue this process until all exposure setting have been tested, except continuous. Compare the series of burns to verify that progressively larger burns occurred as the exposure duration increased.

Repeat Pulse.

If the unit has a repeat pulse feature, which repeats the pulse at a fixed or adjustable rate, test this feature with the laser set at the minimum, median, and maximum repeat pulse settings. Some laser radiometers can react quickly enough to be used to test this feature of the laser. With such a radiometer, test the laser to ensure that the correct power is repeatedly delivered over the correct time period.

If the radiometer cannot be used, use the following alternative test method. Set the laser to about 10 W and a 0.1 sec exposure duration with the fiber, handpiece, or micromanipulator attached, and verify that the Repeat Pulse feature operates as expected by moving the Delrin or colored tongue depressor slightly between each pulse. Be extremely careful to keep hands out of the laser beam path. If the number or duration between repeat pulses is adjustable, test that setting changes made throughout the range result in the expected performance. Verify that pulse output is within 10% or that burns of similar density and size are produced.

Footswitch Exposure Control.

Set the output time for about 5 sec, activate the unit, and release the footswitch after about 1 sec. Verify that the beam turns off when the footswitch is released.

Power Output.

Select one delivery system (e.g., fiber, handpiece, micromanipulator), and perform the manufacturer's recommended user calibration procedure. Secure the delivery system at the appropriate distance from the detector of the laser radiometer to meet spot-size requirements specified in the instructions for the meter. (Do not focus the beam to a small spot on the laser radiometer. Some laser radiometers require that the unfocused or a defocused laser beam be projected into the laser radiometer to cover the majority of the absorber surface. If the laser beam is focused on the detector of such meters, the meter may be damaged.)

WARNING: Accessing the unfocused laser beam may require defeating internal interlocks. Because of the heightened risk associated with an unfocused, nondiverging laser beam, exercise great care if the interlocks are to be defeated.

Note: Laser radiometers use different time constants to acquire an acceptable reading. Identify and use the correct time constant for the radiometer.

Test the laser set at a low (e.g., 10% of full scale), medium (e.g., 50% of full scale), and maximum output. Activate the laser for a sufficient period to acquire acceptable readings. The measured and displayed values should be within 10%. If the laser includes a low-power (e.g., mW) feature, test it in a similar fashion in the radiometer's low-power range. In addition, compare the reading obtained with that taken during acceptance testing, at the last scheduled inspection, or after the last repair.

Laser Delivery System Calibration.

Nd:YAG surgical lasers typically include a user-accessible calibration port leading to an internal radiometer that allows output calibration and testing of the laser fiber. This feature is provided because transmission of laser energy through a fiber can change as a result of fiber use. Based on the measurement from the calibration radiometer, the laser may automatically recalibrate itself and/or adjust displays so that the power indicated to be delivered to the patient will be correct; or it may require the user to do this manually. Verify that this feature is functioning by using the manufacturer's recommended calibration procedure to test one delivery system (e.g., fiber, handpiece) that the manufacturer indicates can be acceptably calibrated using these procedures. (Contact tips cannot be calibrated using the laser's calibration radiometer.)

Preventive Maintenance: General.

Verify that all daily preventive maintenance procedures recommended by the manufacturer are carried out.

Clean.

Clean the exterior. Clean cooling system fibers and accessible optical components (e.g., blast shield, microscope lenses), if necessary, using techniques and cleaning solutions recommended by the manufacturer.

Lubricate.

Lubricate any pump, fan, compressor, or printer components as recommended by the manufacturer.

Calibrate/Adjust.

Calibrate and adjust any components (e.g., printer) according to manufacturer recommendations. Only appropriately trained personnel should attempt laser adjustments. Ensure that all hoses and tubes are tight.

Replace.

Replace filters if needed. Check all fluid levels and supplement or replace fluids if needed.