Argon Surgical Lasers

Procedure No. 462-20081015-01 ( Major )

Used For:

Lasers, Argon, Surgical [16-941]

Commonly Used In:

Operating rooms, short procedure areas, endoscopy laboratories, ophthalmic operating rooms

Scope:

Applies to general-purpose argon surgical lasers that include contact and/or noncontact flexible fiberoptic delivery systems (either reusable or disposable), emit blue-green visible light energy at 514 and 488 nm, and can provide sufficient power output to coagulate and vaporize tissue; applies to low- and high-power argon surgical lasers that are typically used for general surgery, gastroenterology, bronchopulmonary, neurosurgery, gynecology, and ENT surgery procedures; does not apply to ophthalmic argon lasers, which are typically low power (e.g., below 2 W); however, many of the tasks listed herein can be used or modified for these other lasers

Risk Level: High

| Type | Interval | Time Required |
| --- | --- | --- |
| Major | 12 Months | 0 hours |
| Minor | 6 Months | 0 hours |

Overview:

Argon surgical lasers affect tissue by delivering blue-green visible light energy at a sufficient power density to cause vaporization and/or coagulation. The 488/514 nm argon energy is preferentially absorbed by pigmented tissue and hemoglobin and is typically absorbed within 3 mm of the tissue surface. Argon surgical laser fibers are most often used in contact with or close to tissue to cause coagulation and vaporization. Moving the fiber tip away from the tissue to lower the power density causes less tissue to be vaporized and coagulated.

General-purpose argon surgical lasers have a laser tube containing an argon gas mixture that is caused to emit light energy by an electric field. This energy leaves the laser tube through a partially reflecting mirror and is typically directed into a flexible optical fiber that transmits the laser energy to the tissue. The fiber may be used with additional devices (e.g., through an endoscope), with a headpiece and lens, and/or with a laser handpiece or a laser micromanipulator (used to interface the laser with the surgical microscope). These attachments may focus the energy into a small spot size at a known working distance and/or a specific beam direction to accomplish special tasks (e.g., energy can be emitted from the surgeon's headpiece through a handheld lens and focused on the patient's retina). In addition, argon lasers can emit a single pulse or a train of pulses.

Like most lasers, argon lasers are inefficient in converting electrical energy into laser energy. As a result, excess heat is generated in the laser cavity, requiring a cooling system. Most argon lasers use water/air cooling systems that are self-contained, connected to a freestanding chiller system, or connected to a water supply and drain.

Failure of an argon surgical laser can cause patient or staff injury, abrupt interruption of a surgical procedure, or damage to the laser system. These lasers must be meticulously maintained in order to ensure proper and safe operation.

Argon lasers are normally checked before each use by the laser's power-on self-test and by user examination of the aiming beam and the delivery system to be used. This minimizes the need for frequent additional periodic testing. The extent and frequency of inspection by hospital personnel should be coordinated with maintenance by clinical engineering personnel.

Test Apparatus, Supplies, Parts:

- Electrical safety analyzer

- Laser radiometer (0.01 to 20 W, 0.1 to 100 W, 0.01 to 10 sec)

- Calibrated flowmeter (optional)

- Vise with padded jaws or ring stand with padded clamp

- Outlet test fixture (optional)

- Grounding strap (optional)

Special Precautions:

Inspecting and maintaining lasers is a dangerous as well as necessary process, and far greater care is required than with most devices. Personnel who inspect or service lasers should receive special training from the manufacturer or from a qualified alternative training source.

Laser energy can cause serious injury, particularly when the internal interlock is overridden or in any other situation in which the energy does not diverge significantly over long distances. Under some circumstances, the beam may not diverge significantly, even a full room length or more away from the laser (and can harm tissue or burn material even at this distance). Therefore, exercise great care whenever a laser beam is accessible. Area security and use of personnel protective devices and practices should be consistent with hospitalwide laser safety procedures and/or should be approved by the laser safety committee.

In addition, windows should be covered with laser opaque nonreflective material to prevent transmission of laser energy to other areas.

Wear appropriate laser safety eyewear at all times whenever the laser is in the Operating mode. WARNING: Do not stare directly into the aiming system beam or the therapeutic laser beam, even when wearing laser safety eyewear. Avoid placing the laser beam path at eye level (i.e., kneeling, sitting, or standing).

Do not perform these procedures when a patient is present or clinical staff are working, and do not aim the laser across a path that a person might normally use as a thoroughfare. Furthermore, at a minimum, post doors to the room with appropriate laser safety signs stating that the laser is in use and that it is unsafe to enter the room without authorization by the service person performing the procedure. A second person should be present, especially during procedures of recognized risk, to summon help in case of an accident.

The laser should remain in the off position when not in use. When in use, it should be in the standby/disabled mode. Do not switch it to the operating mode until the procedure is about to begin and the laser and its delivery system are properly positioned. If the procedure must be interrupted, disconnect the laser from line voltage, and remove the laser operation key and store it in a controlled location.

Do not use the laser in the presence of flammable anesthetics or other volatile substances or materials (e.g., alcohol), or in oxygen-rich atmospheres, because of the serious risk of explosion and fire. Remove from the working area or cover with flame-resistant opaque material all reflective surfaces likely to be contacted by the laser beam. Whenever possible, use a firebrick or other nonflammable material behind the target material (e.g., black Delrin) when the laser is to be activated. A CO2 fire extinguisher should be readily available.

Some surgical lasers use high voltages (e.g., 20 kV), which can be lethal. Capacitors may store charges long after the device has been disconnected from line voltage. Consult the manufacturer's recommended procedures for servicing high-voltage laser circuits, and avoid contact with any portion of the high-voltage circuit until you are certain that the charge has been drained. In such instances, a good ground must be present; preferably, use a redundant ground strap if you must enter the laser cabinet. When possible, disconnect the laser from line voltage before entering the laser cabinet, and use insulated gloves for those procedures in which contact with a high-voltage source is possible (and the gloves are not otherwise contraindicated). Ensure that equipment intended to be used to measure, drain, or insulate high voltages carries the appropriate insulation rating (e.g., above 20 kV).

Where possible, perform tests with the unit turned off. Because of the presence of high voltage, perform the Grounding Resistance test before any other test that requires operation of the laser.

WARNING: Lasers may be damaged by switching between normal and reverse polarity while the device is on. If reverse-polarity leakage current measurements are made, turn off the unit being tested before switching polarity. Also, lasers powered by three-phase electrical systems may be damaged if proper electrical phase connections are not made initially and maintained thereafter. Thus, do not conduct  reversed-polarity leakage current testing on three-phase powered lasers.

WARNING: Do not use an anesthesia or other similar bag that may have a mold-release agent (e.g., starch, talc) on its inside surface because the agent could contaminate the gas recirculation system of the laser and ultimately contaminate a patient wound during a subsequent procedure.

Report any laser accident immediately to the laser safety officer or equivalent, as well as to the hospital risk manager.

Procedure:

Be sure that you understand how to operate the equipment, the significance of each control and indicator, and the alarm capabilities. Before beginning an inspection, carefully read this procedure, the operators manual, and the inspection and preventive maintenance procedures recommended by the manufacturer (typically included in the service manual). Use the BiomedicalBenchmark Support Assessment Form to document a maintenance decision that reflects past experience with this type of equipment and the environment where it is used. Then use the IPM Procedure Customization Tool to modify this procedure as needed; the program will generate a documentation form with the corresponding changes.

Qualitative tasks: Chassis/Housing.

Verify that the key has not been left in the laser. (Remove it if it has been, and inform users of the importance of storing the key in a controlled location.) Examine any external gas tanks that may be in use with the laser, and ensure that they have been turned off after the last use. Examine the exterior of the unit for cleanliness and general physical condition. Be sure that all housings are intact and properly aligned, that assembly hardware is present and tight, that any retractable parts slide easily and lock in place if so constructed, that there are no signs of spilled liquids or other evidence of abuse, and that there are no obvious signs of water or oil leakage.

Shutters.

If manual shutters for the aiming system or the therapeutic lasers are accessible, ensure that they operate smoothly and correctly. Be sure to leave the shutter in the proper position for normal operation.

AC Plug.

Examine the AC power plug for damage. Wiggle the blades to determine whether they are secure. Shake the plug, and listen for rattles that could indicate loose screws. If damage is suspected, open the plug and inspect it.

Line Cords.

Inspect line cords for signs of damage. If a cord is damaged, replace the entire cord, or, if the damage is near one end, cut out the defective portion. Be sure to wire a new power cord or plug with the correct polarity.

Strain Reliefs.

Examine the strain reliefs at both ends of the line cord. Be sure that they grip the cord securely.

Tubes/Hoses.

Check the condition of all cooling-system hoses and any other hoses or tubing the laser may have (e.g., drain, gas). Check that they are of the correct type; that they have not become cracked and do not show other signs of significant abuse; that they are connected correctly and positioned so that they will not leak, kink, trail on the floor, or be caught in moving parts; and that they are secured adequately to any connectors.

Cables.

Inspect all cables and their channels or strain reliefs for general physical condition. Examine cables carefully to detect breaks in insulation and to ensure that they are gripped securely in the connectors at each end to prevent strain on the cable.

Fittings/Connectors.

Examine all optical (e.g., fiber), gas, liquid, and electrical fittings and connectors for general physical condition. Gas and liquid fittings should be tight and should not leak. Electrical contacts should be straight, clean, and bright.

If external gas tanks or wall-supply outlets can be used, gas-specific connectors should be present. Be sure that no pins are missing from yokes and that the keying and indexing of connectors for each gas to be used is correct. A laser that connects to a central piped medical gas system or to a freestanding medical gas system should have the matching DISS or quick-connect fitting for the gas that it is to be used with. Verify that suitable unique connectors are supplied so that adapters are not required.

Filters.

Check the condition of all liquid and air filters. Some argon surgical lasers require deionized water, and most require special filtration. Measuring the pressure drop across a liquid filter can be helpful in determining whether the filter should be replaced. Clean or replace filters according to manufacturer recommendations (e.g., replace if the pressure drop is >5 psi), and indicate this in the preventive maintenance section of the inspection form. Clean or replace air filters that are obviously dirty. Verify operation of safety filters in the microscope and endoscope delivery systems.

Controls/Switches.

Examine all controls and switches for physical condition, secure mounting, and correct motion. If a control has fixed-limit stops, check for proper alignment as well as positive stopping. Check membrane switches for tape residue and for membrane damage (e.g., from fingernails, pens, or surgical instruments). If you find such evidence, notify users to avoid using tape and sharp instruments. During the inspection, be sure that each control and switch works properly.

Remote Control.

Examine the exterior of the control for cleanliness and general physical condition. Be sure that housings are intact, that assembly hardware is present and tight, and that there are no signs of spilled liquids or other serious abuse. If the remote control is attached by cable to the laser, ensure that the cable and any connectors are in good condition. Examine all controls and switches for general physical condition, secure mounting, correct motion, and intended range of settings. Where a control should operate against fixed-limit stops, check for proper alignment as well as positive stopping. During the course of the inspection, be sure to check that each control and switch performs properly.

Footswitch.

Examine the footswitch for general physical condition, including evidence of spilled liquids. Footswitches for lasers include an internal switch that activates according to the depth of pedal depression. It is usually possible to feel the vibration caused by closure of the switch, even through a shoe. Check that the internal switch is operating and that the footswitch does not stick in the on position. Some footswitches include two internal switches; in this case, verify the operation of both. Some footswitches also include a switch to operate the liquid- or gas-cooling system. Check to be sure that this switch operates reliably.

During the procedure, check to be sure that the laser activates consistently when the footswitch is depressed and that the fiber-coolant system operates properly when the fiber-coolant switch is activated and deactivated. Flex the cable at the entry to the switch, and, using an ohmmeter, check for internal wire breaks that cause intermittent operation. Confirm that strain reliefs are secure.

Examine the male and female connectors for attaching the footswitch to the laser cabinet to be sure that no pins are bent and that no other damage is present. Ensure that the connector secures acceptably to the laser cabinet.

Microscope Micromanipulator.

Examine the microscope micromanipulator for cleanliness and general physical condition. Be sure to handle it by the main body; do not hold it by the joystick, and do not touch the reflecting lenses in the body. Inspect micromanipulators provided by both the laser manufacturer and the laser accessory manufacturer.

Ensure that the reflecting lenses are intact and clean. Consult the manufacturer's recommendations for the procedures and cleaning agents to use to clean reflecting lenses.

Examine the joystick to ensure that it is firmly attached and that it freely moves the reflecting lens. If a finger rest is present, ensure that it is firmly attached and properly oriented.

If a zoom focus feature is present, be sure that it turns easily and does not slip. Examine each objective lens to ensure that it is intact and clean. Do not touch the lens surface. Consult the manufacturer's recommendations for the procedures and cleaning agents to use to clean the objective lenses. Carefully insert each lens into the micromanipulator, and ensure that it fits snugly.

There should be no visible dirt or residue in the optical path of the laser aperture. Ensure that any mechanism to close off the laser aperture (fiber port) is clean, operates smoothly, and is in use.

Inspect the mechanism used to attach the micromanipulator to the microscope and to the laser to ensure that all parts are present and that it is in good working order. Connect the micromanipulator to the microscope and to the laser to check for a secure connection.

Pumps/Fans/Compressors.

Check the physical condition and proper operation of these components, if present. If lubrication is required, note this in the preventive maintenance section of the form.

Fluid Levels.

Check all fluid (e.g., coolant) levels. Refill or change the fluid according to the manufacturer's recommendations, and note this in the preventive maintenance section of the inspection form. If an external water supply is in use, ensure that the water pressure is properly regulated and at the appropriate pressure and that the supply and drain system is properly configured (e.g., filters are oriented for proper flow, drain hoses are positioned in a sink or drain).

Battery.

Inspect the physical condition of batteries and battery connectors, if readily accessible. If a remote control or display is battery powered, check or replace the battery (periodic prophylactic battery replacement is often preferred to risking battery failure during use). When it is necessary to replace a battery, label it with the date.

Indicators/Displays.

During the course of the inspection, verify proper operation of all lights, indicators, meters, gauges, and visual displays on the unit and remote control. Ensure that all segments of a digital display function. Note any error messages displayed during the power-on self-test.

If primary and remote-control indicators and displays can be used at the same time or if control can be switched from one to the other during the course of a procedure, verify that the same information (e.g., settings, displays) is indicated on both control panels during laser operation.

If display screens or digital displays are provided for user prompts or for viewing accumulated information (e.g., pulse or accumulated energy counter), ensure that each display provides the information expected. Ensure that user prompts occur in the proper sequence. Store some sample information, and verify that it is correct. If a feature to manually reset this information is available, ensure that it works.

Time/Date Settings.

Verify that the time and date settings on the unit are correct.

Alarms/Interlocks.

Operate the device in a manner that will activate the self-check feature, if present, and verify that all visual and audible alarms activate according to the manufacturer's documentation. If no self-check feature is present, operate the laser in a manner that will activate each audible and visual alarm; be sure to test only those alarms that will not cause damage to the laser or present an unnecessary risk of laser beam exposure to the user or bystanders.

If a door or window interlock is used, ensure that it deactivates the laser properly. (Do not disassemble major parts of the laser to test internal interlocks.) After deactivating the laser and reclosing the door or window, check to be sure that the laser will restart. Be sure to check the interlocks in all locations where the laser is used. (For some lasers, the function of the interlocks can be checked using an ohmmeter.)

If the laser is equipped with an emergency "kill" switch, test this feature to be sure that it deactivates the laser and that the laser will subsequently restart.

Audible Signals.

Operate the device to activate any audible signals (e.g., laser emission, setting change). Check for proper operation, and verify that the signal can be heard in the environment in which the laser will be used.

Labeling.

Check that all placards, labels, and instruction cards noted during acceptance testing are present and legible. Check to see that an instruction manual is kept with the laser or is readily available.

Accessories.

Verify that all necessary accessories are available and in good physical condition. Set up reusable accessories with the laser to ensure compatibility and proper functioning.

Checking all fibers or accessories during a single inspection and preventive maintenance procedure is unnecessary as long as accessories are routinely checked by the person(s) responsible for laser setup and operation. In addition, many of the accessories are sterile and would require resterilization before use, making the laser potentially unavailable. Be sure to check with the person responsible for scheduling the use of the laser before beginning the procedure.

Fibers.

For the test fiber or before each use, examine the connector, cable, and tip of each fiber that may be used, as well as the fiber support, for cleanliness and general physical condition. Be sure that all hardware (e.g., laser gas tubing channels) is present, in good condition, and firmly attached. Ensure that the connector properly seats into the laser aperture of the laser cabinet. Examine the distal end of fibers to ensure that any connecting mechanisms (e.g., threads) are in proper working order.

If a fiber appears to be dirty or damaged, remove it from service. If a fiber is reusable, notify the person(s) responsible for fiber repair. The fiber should be repaired and/or cleaned according to the manufacturer's recommendations. Verify fiber performance.

Handpieces.

Examine each handpiece component (e.g., body, tips, lenses) for cleanliness and general physical condition. Examine individually only those components that are intended for removal during normal use and storage. (Do not remove other parts that are press-fit or attached by screws, bolts, or snap-rings.) If lenses are detachable, be sure not to touch the lens surface; handle lenses by the edges only. Consult the manufacturer's recommendations for the procedures and cleaning agents to use to clean lenses.

Ensure that major subcomponents of the handpiece, when assembled, are secure. Ensure that the mechanisms used to connect the handpiece(s) to the fiber are in good working order and that they reliably secure each handpiece to the fiber.

Aiming Beam.

Argon lasers typically use an attenuated therapeutic beam as the aiming beam. Activate the aiming beam (without the therapeutic beam), and verify that it produces a round, uniformly bright spot, with no halo. For handpieces that provide adjustable spot sizes, verify that the spot size changes as expected and still remains uniform.

Laser Aperture.

WARNING: Make this inspection with the laser powered off.Remove and inspect the protective window (e.g., blast shield) behind the laser aperture. It should be clean and undamaged; replace if needed. There should be no visible dirt or residue in the optical path of the laser aperture.

Gas Regulators.

Examine any gas regulators for cleanliness and general physical condition. Ensure that the gauges on the regulators are not broken. During the procedure, ensure that the regulator and the gauge operate as expected. Verify that the correct gas is attached to each regulator. Be sure that a key or wrench to facilitate changing the gas supply is with the unit or readily available.

If the laser includes a gas recirculation system, ensure proper operation by allowing it to control the gas supply into and out of a sealed plastic bag.

WARNING: Do not use an anesthesia or other similar bag that may have a mold-release agent (e.g., starch, talc) on its inside surface because the agent could contaminate the gas recirculation system of the laser and ultimately contaminate a patient wound during a subsequent procedure. If proper operation is questionable, consider using a calibrated flowmeter to measure actual gas flow.

Quantitative tasks: Grounding Resistance.

Measure the resistance between the grounding pin of the power cord (if so equipped) and exposed (unpainted and not anodized) metal on the chassis. Grounding resistance should not exceed 0.5 Ω. If the unit is double insulated, grounding resistance need not be measured.

Exposure Duration.

Some laser radiometers can measure pulse duration. If the power meter can react to pulse duration (this is the preferred circumstance), test the laser at each setting. However, if the radiometer does not measure pulse duration, use the following less preferable alternative.

Place and secure the laser fiber, handpiece, or micromanipulator with the aiming system focused on the target material (e.g., black Delrin or a tongue depressor). With the laser set to 10 W and the exposure set at minimum duration, activate the laser and create a burn. Carefully move the target material to expose a clean area, maintaining the same distance. Adjust the exposure setting in increments of 0.1 sec or the next longest duration, and activate the laser at each setting. Continue this process until all exposure setting have been tested, except continuous. Compare the series of burns to verify that progressively larger burns occurred as the exposure duration increased.

Repeat Pulse.

If the unit has a repeat pulse feature, which repeats the pulse at a fixed or adjustable rate, test this feature with the laser set at the minimum, median, and maximum repeat pulse settings. Some laser radiometers can react quickly enough to be used to test this feature of the laser. With such a radiometer, test the laser to ensure that the correct power is repeatedly delivered over the correct time period.

If the radiometer cannot be used, use the following alternative test method. Set the laser to about 10 W and a 0.1 sec exposure duration with the fiber, handpiece, or micromanipulator attached, and verify that the Repeat Pulse feature operates as expected by moving the Delrin or colored tongue depressor slightly between each pulse. Be extremely careful to keep hands out of the laser beam path. If the number or duration between repeat pulses is adjustable, test that setting changes made throughout the range result in the expected performance. Verify that pulse output is within 10% or that burns of similar density and size are produced.

Footswitch Exposure Control.

Set the output time for about 5 sec, activate the unit, and release the footswitch after about 1 sec. Verify that the beam turns off when the footswitch is released.

Power Output.

Select one delivery system (e.g., fiber, handpiece, micromanipulator), and perform the manufacturer's recommended user calibration procedure. Secure the delivery system at the appropriate distance from the detector of the laser radiometer to meet spot-size requirements specified in the instructions for the meter. (Do not focus the beam to a small spot on the laser radiometer. Some laser radiometers require that the unfocused or a defocused laser beam be projected into the laser radiometer to cover the majority of the absorber surface. If the laser beam is focused on the detector of such meters, the meter may be damaged.)

WARNING: Accessing the unfocused laser beam may require defeating internal interlocks. Because of the heightened risk associated with an unfocused, nondiverging laser beam, exercise great care if the interlocks are to be defeated.

Note: Laser radiometers use different time constants to acquire an acceptable reading. Identify and use the correct time constant for the radiometer.

Test the laser set at a low (e.g., 10% of full scale), medium (e.g., 50% of full scale), and maximum output. Activate the laser for a sufficient period to acquire acceptable readings. The measured and displayed values should be within 10%. If the laser includes a low-power (e.g., mW) feature, test it in a similar fashion in the radiometer's low-power range. In addition, compare the reading obtained with that taken during acceptance testing, at the last scheduled inspection, or after the last repair.

Laser Delivery System Calibration.

Some argon surgical lasers include a user-accessible calibration port or radiometer that allows output calibration and/or testing of the laser fiber. This feature is provided because transmission of laser energy through a fiber may change as a result of fiber use. Based on the measurement from the calibration radiometer, the laser may automatically recalibrate itself and/or adjust the displays so that the power indicated to be delivered to the patient will be correct; or it may require the user to do this manually. Verify that this feature is functioning by using the manufacturer's recommended calibration procedure to test one delivery system (e.g., fiber, handpiece) that the manufacturer indicates can be acceptably calibrated using these procedures. A good quality (e.g., >85% transmissibility, undamaged sheath) fiber or handpiece should be used for this test.

Preventive Maintenance: General.

Verify that all daily preventive maintenance procedures recommended by the manufacturer are carried out.

Clean .

Clean accessible optical components (e.g., blast shield, microscope lenses), if necessary, using techniques and cleaning solutions recommended by the manufacturer.

Lubricate.

Lubricate any motor, pump, fan, compressor or printer components as recommended by the manufacturer.

Calibrate/Adjust.

Calibrate and adjust any components (e.g., printer) according to the manufacturer recommendations. Only appropriately trained personnel should attempt laser adjustments. Ensure that all hoses and tubes are tight.

Replace.

Replace filters if needed. Check all fluid levels and supplement or replace fluids if needed.