Syringe Pumps

Procedure No. 488-20081015-01 ( Major )

Used For:

Drivers, Syringe [18-092]

Infusion Pumps, Syringe  [13-217]

Commonly Used In:

Operating rooms, labor and delivery rooms, NICUs, other patient care areas

Scope:

Applies to electronic syringe pumps and mechanical syringe drivers

Risk Level: High

| Type | Interval | Time Required |
| --- | --- | --- |
| Major | 12 Months | 0 hours |
| Minor | 0 NA | 0 hours |

Overview:

Syringe pumps are often used to deliver small volumes of medications and solutions sometimes over long periods of time. Their flow range is usually 0.1 mL/hr – 99.9 mL/hr. Syringe pumps and drivers operate by securing a syringe barrel and advancing its plunger. Syringe pumps provide very accurate delivery using a stepper motor and lead screw. Syringe drivers provide less resolution in selecting flow and are less accurate in delivery as they are typically spring-powered.

Test Apparatus, Supplies, Parts:

- Electrical safety analyzer

- Pressure meter (≥50 psi)

- Infusion analyzer optional

- 50 mL graduated cylinder (1 mL graduations)

- Stopwatch or watch that displays seconds

- Electronic balance (200 g range, 0.1 mg resolution)

- 5 mL pipette (0.1 mL graduations)

Special Precautions:

Procedure:

Be sure that you understand how to operate the equipment, the significance of each control and indicator, and the alarm capabilities. Before beginning an inspection, carefully read this procedure, the operators manual, and the inspection and preventive maintenance procedures recommended by the manufacturer (typically included in the service manual). Use the BiomedicalBenchmark Support Assessment Form to document a maintenance decision that reflects past experience with this type of equipment and the environment where it is used. Then use the IPM Procedure Customization Tool to modify this procedure as needed; the program will generate a documentation form with the corresponding changes.

Use the disposable accessory (e.g., syringe and possibly tubing set) specified for the pump to be inspected. Disposable accessories should be replaced at least weekly and more frequently with heavy use.

For settings ≥1 mL/hr, flow accuracy is most conveniently determined with an infusion pump tester. If a pump tester is not available, calculate flow rate by dividing the volume collected in a 50 mL graduated cylinder (1 mL graduations) by the delivery time measured with a stopwatch. For settings less than 1 mL/hr, calculate flow rate by dividing the mass obtained in a small beaker on an electronic balance (200 mg range and resolution to 0.1 mg) by the measurement time. If a balance is not available, use a 5 mL pipette (0.1 mL graduations) attached to a vertical mounting stand and divide the volume collected by the delivery time.

Operate the pump on battery power during during a flow accuracy measurement if not during the entire inspection to check that the battery has been charged and can hold a charge. If a low-battery alarm occurs, verify that the alarm is properly displayed and then continue the inspection on line power. Note how long the pump was operating on battery power and the conditions under which the low-battery alarm occurred. Fully charge the battery before returning the pump to use.

Qualitative tasks: Chassis/Housing.

Examine the infusion device for overall condition. The chassis should be clean and free from IV or enteral solution residue, especially near moving parts (e.g., thumbwheel switches, pump or controller mechanisms). Also check for dried solution deposits on accessible air-in-line sensors, pressure sensing mechanisms, and infusion set/cassette locking mechanisms. Check that labels and markings are legible.

Mount.

Screws and brackets that attach the infusion device to an IV pole should be secure and functioning. If the device is mounted on a stand or cart, examine the condition of the mount. Also examine the pole, stand, or cart.

AC Plug.

Examine the AC power plug for damage. Attempt to wiggle the blades to determine that they are secure. Shake the plug and listen for rattles that could indicate loose screws. If any damage is suspected, open the plug and inspect it.

If the device or its IV pole has electrical receptacles for accessories, inspect them by inserting an AC plug into each and checking that it is held firmly. If accessories are plugged and unplugged often, consider a full inspection of the receptacle.

Line Cord.

Inspect the cord for signs of damage. If damaged, either replace the entire cord or, if the damage is near one end, cut out the defective portion.

Strain Reliefs.

Examine the strain reliefs at both ends of the line cord. Be sure that they hold the cord securely.

Fittings/Connectors.

Examine any electrical cable connectors for general condition. Electrical contact pins or surfaces should be straight and clean. Check any spill-protection connector caps for signs of damage.

Controls/Switches .

Examine all controls and switches for physical condition, secure mounting, and correct motion. Where a control should operate against fixed-limit stops, check for proper alignment, as well as positive stopping. Check membrane switches for membrane damage (e.g., from fingernails, pens). During the course of the inspection, be sure to check that each control and switch performs its proper function.

Battery.

If so equipped, check the pump's battery log for battery status. Operate the pump on battery power during during a flow accuracy measurement if not during the entire inspection to check that the battery has been charged and can hold a charge. If a low-battery alarm occurs, check to ensure that it is properly displayed and then continue the inspection using line power. Note how long the infusion pump has been operating and the conditions under which the low-battery alarm occurred. Fully charge the battery before returning the pump to use. When it is necessary to replace a battery, label it with the date.

Indicators/Displays.

During the course of the inspection, confirm the operation of all lights, indicators, meters, gauges, visual displays, and display backlighting, if so equipped. Be sure that all segments of a digital function. (Many infusion devices automatically check indicator and display function when turned on or during a manually activated self-test.)

Self-Test.

For units with a self-test mode, activate it and determine if the expected response is produced.

Time/Date Settings.

Verify that the time and date settings on the unit are correct.

Network/Wireless Interfaces.

Review measures taken to ensure protection against the potential risks and vulnerabilities to the confidentiality, integrity, and availability of electronic information stored or transmitted by the device or system and verify that preventive measures are still active. For example, are passwords being applied correctly, have operating system and virus protection patches and upgrades been installed, is the device still operating on a VPN (virtual private network) and are wireless security measures still in place. Verify that data backup processes are activated and that data can be retrieved from backups.

Alarms.

Many infusion device alarm capabilities can be checked qualitatively. The following procedures include tests for the most common alarm conditions. Check the instruction manual to see how the alarm should work. When an alarm occurs, check to see that both audible and visual alarms are activated and that flow stops or is reduced to a keep-vein-open rate (e.g., <5 mL/hr). Confirm appropriate alarm volume, as well as the operation of any volume control.

Audible Signals.

Operate the device (e.g., press rate switches) to activate any audible signals. Confirm appropriate volume, as well as the operation of the volume control, if so equipped.

Misloaded Syringe.

Check this alarm during setup and operation.

Empty Syringe.

Load a syringe with its plunger completely depressed and verify that the pump alarms for and indicates end of infusion.

Occlusion.

Verify operation of downstream occlusion detection by activating infusion with the infusion set's distal connector capped.

Nurse Call (verify only if this function is used).

Some pumps have a relay contact closure that activates a nurse-call system when an alarm condition occurs. This requires a special cable that connects the pump to the nurse-call system. If the infusion device has this capability and it is used in any clinical location, connect the cable, and simulate one or more of the above alarm conditions to determine whether they activate the nurse call. Alternatively, use an ohmmeter to check that a change in resistance (either low to high or high to low) occurs between the two conductors of the cable when an alarm condition is created.

Labeling.

Check that all necessary placards, labels, and instruction cards are present and legible.

Quantitative tasks: Grounding Resistance.

Measure the resistance between the grounding pin of the power cord (if so equipped) and exposed (unpainted and not anodized) metal on the chassis. Grounding resistance should not exceed 0.5 Ω. If the unit is double insulated, grounding resistance need not be measured.

Flow Accuracy.

Determine the flow accuracy at 10.0 and 100 (99.9) mL/hr.

Use an infusion pump tester or collect the output in a graduated cylinder. Use a stopwatch or a watch that displays seconds to time the delivery into the graduated cylinder until at least 10 mL is collected. Record the time interval and volume collected, and calculate the delivery rate in mL/hr.

To calculate flow error, use the following formula:

Most syringe pumps are specified to deliver within 5% of the flow setting. Note: Negative and positive flow error represents underdelivery and overdelivery, respectively.

Maximum Pressure.

Reference the pump's maximum pressure determined during acceptance testing. Connect the distal end of the primed administration set to a pressure meter or use an infusion pump tester (if equipped to perform this task). Start infusion at a commonly used flow setting (e.g., 10 or 100 mL/hr) and record the maximum pressure. Restart the pump after an occlusion alarm to ensure that maximum infusion pressure has been attained.

The maximum pressure of most pumps is less than 20 psi; it should be within 1 psi of manufacturer specification. Small syringes may produce pressures that exceed 20 psi.

Preventive Maintenance: Replace.

Replace the primary battery and/or the clock/memory battery/cell, if necessary, and label it with the date.